## Disclaimer
2013-06-01
The NCS frequently publishes reports for fellow professionals in which recommendations are given for
various quality control procedures or otherwise. The members of the NCS board and the members of
the concerning subcommittee do not claim any authority exceeding that of their professional expertise.
Responsibility on how the NCS recommendations are implemented lies with the user, taking into
account the practice in his/her institution.
This spreadsheet is to be used as a starting point, not as an exhaustivesource.

## PRI
STEP 3 | STEP 4-A | STEP 4-E | STEP 4-F |  | STEP 4-B |  |  | STEP 4-C | STEP 4-D |  |  | STEP 4-E | STAP 5
Discission tree
1 | 2 | 3 |  |  | 4 | 5 |  | 6 | 7 |  |  | 8 | 9 | 10 | 11 | 12 | 13
Eliminate,
[5 empty] | Risk | [7 empty] | control |  |  | Resonsible | Management
Proces step | Potential Failure Modes | Potential Causes | Potential Consequences | Reference | Frequency | Severity | Undetectability | RPN (F.S) | Critical step | Controlability | Detectability | Continue | or accept this risk | Description of action taken | Possible result | person | approval
http://www.who.int/patientsafety/activities/technical/radiotherapy_risk_profile.pdf
Patient selection
choice of treatment | designated as IMRT patient but does not fulfill IMRT criteria | interpretaion/human error | treatment delay |  | 3 | 1 | 2 | 3 |  |  | yes
| not designated as IMRT patient but does fulfill IMRT criteria |  | higher incidence of complications |  | 2 | 3 | 3 | 6 |  |  | yes
info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Imaging and preparation
Primary image set - CT | faulty calibration of CT scanner | protocol violation during calibration | Incorrect dose distribution |  | 1 | 4 | 9 | 4 |  | yes |  |  |  | Perform regular QA
| wrong position (HFS - FFS) | deviation from protocol, and not mentioned in notes | treatment delay |  | 2 | 3 | 3 | 6 | yes |  |  | yes | control | Check scans at import
| no tattoos placed |  | treatment delay |  | 2 | 2 | 2 | 4 |  |  | yes
Primary image set - CT for Boost
Secondary image set - MRI, PET-CT, CT
Positioning aids | incorrect aids used | deviation from protocol, and not mentioned in notes | treatment delay |  | 3 | 3 | 2 | 9 |  |  |  | yes | Control | Add checks
|  |  | sub-optimal treatment |  | 3 | 3 | 8 | 9 |  |  |  | yes | Control | Add checks
Matching of primary with secondary image sets | match of primary and secondary image set are not accurate | MRI used for delineation of soft tissue - no geometrical 3D correction applied for high gradient distortions | sub-optimal treatment |  | 3 | 3 | 7 | 9 |  |  |  | yes | Control | Adjust scan protocols, add checks while importing secondary images
|  | any modality - Secondary image set is not acquired according to protocol or mentioned in notes from imaging | incorrect treatment plan |  | 2 | 2 | 5 | 4
|  |  | treatment delay due to rescan |  | 3 | 2 | 2 | 6
|  | MRI unexpected gradient distortions | sub-optimal treatment
|  | CT external - different HU | sub-optimal treatment |  | 1 | 2 | 10 | 2
| match is faulty, mix-up of orientations (L-R, Cra-Cau) | incorrect patient positioning (FFS instead of HFS) | treatment delay due to rescan |  | 2 | 4 | 1 | 8 |  |  |  |  | Control | Add checks while importing secondary images
incorrect target definition
| match cannot be performed | any - Image set is not acquired correctly, i.e. according to protocol | treatment delay due to rescan |  | 3 | 2 | 2 | 6
|  |  | treatment is less accurate |  | 3 | 2 | 2 | 6
Organ delineation | erronous delineation of PTV (mix-up) | too many organs delineated | incorrect treatment plan |  | 2 | 3 | 4 | 6
| erronous delineation of GTV | faulty delineation | incorrect treatment plan
| incorrect delineation of organ | delineation not according to protocol (rectum too short) | incorrect treatment plan |  | 2 | 3 | 7 | 6
|  | expanded structure not cleared, thereby faulty steering the optimisation | incorrect treatment plan |  | 2 | 2 | 9 | 4
| non standard organ (not in script) is not delineated |  | higher incidence of complication
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Scheduling Patient
Treatment Planning
Importing image sets (primary and Secondary) | wrong image set imported | many image sets to choose from | incorrect treatment plan |  | 2 | 3 | 7 | 6 |  |  |  |  | Control
Delineation of Organs | delineation CTV, GTV
| incorrect margings applied to GTV to make PTV | unclear protocol or human error | increased toxicity and decreased tumour control |  | 2 | 3 | 9 | 6 |  |  |  |  | Control | Include GTV/CTV/PTV expansions of delineated structures in macro's/scripts. Record imporant planning activities thast relate to GTV/CVT/PTV definition to allow for automatic evaluation of followed procedures post planning.
| during expansion of contours artifacts are created, and not cleared by the operator | software feature / protocol violation | sub-optimal treatment plan |  | 2 | 3 | 7 | 6 |  |  |  |  | Eliminate / Control | Software update / additional check by planner
Copy PTV's | incorrect comination of PTV structures that are used for optimization (PTV_glands, PTV_surdo, PTV_primiry etc) | Unclear which PTV's should be summed or subtracted | Incorrect definition of PTV and thus over or underdosage |  | 2 | 3 | 7 | 6
Selection of fractionation scheme | wrong choice of SIB / Sequential Boost | the fractionation scheme is selected by the MD, based on several test dose distribution. | sub-optimal treatment plan |  | 2 | 3 | 9 | 6
wrong fractionation schedule entered
Placement of beams | incorrects choice of beam angles | incorrect macro/script used | sub-optimal treatment plan |  | 1 | 3 | 9 | 3
|  | class solution is not optimal for this patient | sub-optimal treatment plan |  | 3 | 2 | 7 | 6
Enter dose constraints/objectives | incorrect choice of objecties / constraints | incorrect macro/script used | sub-optimal treatment plan |  | 3 | 3 | 9 | 9 |  |  |  | yes | Eliminate / Control | Use marco/script to generate a list of dose constraints/objectives. Add check for used constraints/objectives.
|  | class solution is not optimal for current patient | sub-optimal treatment plan |  | 1 | 3 | 9 | 3
|  | incorrect macro/script used | non accurate dose distributions |  | 2 | 2 | 9 | 4
Placement dose grid | incorrect dose grid | incorrect entry of grid resolution | non accurate dose distributions | Chung Jin Palta | 2 | 2 | 6 | 4
|  | incorrect entry of grid size | non accurate/incorrect dose distributions |  | 2 | 3 | 6 | 6
Optimization / inverse planning | unrealistic fluency paterns or uncurrent segment shapes | optimisation driven by CT artifacts (markers/shadow) | sub-optimal treatment plan |  | 3 | 2 | 7 | 6
|  | optimisation driven by low CT density areas or built up regions | sub-optimal treatment plan |  | 3 | 2 | 7 | 6
|  | optimisation driven by faulty delineated organs | sub-optimal treatment plan |  | 3 | 3 | 4 | 9 |  |  |  | yes | Control | Add check to guarantee delineation of organs is done according to protocol and/or check soundness of delineated organs
objective demands results in very high dose gradients
|  | poor performance of sequencer | quality of the dose distribution is not guaranteed
| use of normalisation point during optimisation | entry of presciption dosis in ISOC according to ICRU | sub-optimal plan because the solution space is limited by using a normalisation point during optimisation. |  | 4 | 2 | 6 | 8 |  |  |  |  | Eliminate | prohibit this approach of planning
| use of normalisation point after optimisation | entry of presciption dosis in ISOC according to ICRU | sub-optimal treatment plan to catastrophic. The use of a normalisation point conflicts with full optimisation of the dose distribution. |  | 4 | 4 | 6 | 16 |  |  |  | yes | Control | Add check when a normalisation point is used after full optimisation of the dose distribuion
Multiple optimization loops or iterative optimization | non-optimal plan | incorrect use of objectives | sub-optimal treatment plan | Holdsworth Med Phys 2964-; 38 (6) 2011 | 2 | 3 | 9 | 6
|  | violation of optimisation protocol (number of iterations) |  |  | 3 | 2 | 5 | 6
|  | protocol is not available |  |  | 2 | 2 | 6 | 4
|  | use of not-validated paramters and/or biological objectives (EUD) | sub-optimal treatment plan |  | 2 | 3 | 9 | 6
| plan has become complex, undermining the quality of the dose calculation. | optimization performed too often or too long | sub-optimal treatment plan |  | 3 | 2 | 4 | 6
| incorrect optimisation parameters | manner to improve the calculated dose distribution | sub-optimal treatment plan |  | 3 | 2 | 4 | 6
Plan evaluation | evaluation of plan fails | wrong plan chosen | sub-optimal treatment plan |  | 2 | 2 | 10 | 4
|  | mixing up of isodose lines (percentage or absolute) |  |  | 1 | 4 | 4 | 4
|  | evaluation incorrectly carried out |  |  | 3 | 3 | 9 | 9 |  |  |  | yes | Control | Use checklist
|  | plan is not evaluated |  |  | 2 | 2 | 4 | 4
Plan export to R&V | incorrect plan exported to MQ | incorrect plan chosen for export | sub-optimal treatment plan |  | 2 | 3 | 10 | 6 |  |  |  |  | Eliminate | Adjust workinstrunctions: remove obsolete plans during plannning after selection of clinical treatment plan.
| corrupt plan received | network inconsistency | from undeliverable plan to catastrophic |  | 1 | 4 | 6 | 4
| incomplete data set received in MQ, manual entering missing data | not all treatment paramters are send or received | from undeliverable plan to catastrophic |  | 2 | 4 | 2 | 8 |  |  |  |  | Eliminate
Plan control | independent dose check is not possible for IMRT | infrastructure does not support IMRT plan checks | measurements are needed, and may cause treatment delay or unneccecary use of human resources. |  | 2 | 2 | 1 | 4 |  |  | yes
| during plan checks in the R&V system, unintended alterations are entered in the plan | typo | adjustment of treatment plan |  | 1 | 3 | 10 | 3 |  |  |  |  | Eliminate / Control | Adjust software to guarantee that this cannot happen without notification
|  | bad design of software | treatment deviates strongly from intentional plan | [9 empty] | Eliminate | Adjust software to guarantee that this cannot happen without notification
| MU are corrected for couch absorption manually, i.e. outside the TPS | typo (. vs , or none at all) | catastrophic |  | 1 | 4 | 4 | 4 |  |  |  |  | Eliminate / Control | Adjust software to warn if . vs , typo was made. Add check that compares old and new value of MU, using a certain bandwidth.
|  | wrong calculation on calculator | sub-optimal treatment plan to catastrophic |  | 1 | 4 | 10 | 4
| misinterpretation of dose distribution | lines are absolute or relative distributions | sub-optimal treatment plan to catastrophic |  | 1 | 4 | 6 | 4
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
| data ambiguity | multiple data sources may be used independently (patient file, patient information management system, treatment planning system) for total dose, fraction dose, number of fractions, number of fractions per week | sub-optimal treatment plan to catastrophic |  | 4 | 4 | 4 | 16 |  |  |  | yes | Eliminate | Use one complete data set source, only allow for export of data to other systems, no manual data entry
Planning of boost on new CT | inaccurate dose distrubutions of total dose | because patient geometry has changed during treatment a new CT scan was made and used for treatment planning. Plans and dose distributions on different CT scans cannot be added | increased toxicity and/or decreased tumour control because of conservative planning |  | 3 | 3 |  | 9 |  |  |  | yes | Eliminate | Software solutions
Re-planning during treatment | wrong number of fractions used for remainder of treatment | optimisation is performed as if this was a new plan, with the objectieves and fractions. After optimisation number of fractions should be made in accordans with clinic, but is forgotten. More fractions delivered | sub-optimal treatment plan to catastrophic |  | 1 | 4 | 1 | 4 |  |  |  |  | Control | Add checks
| correct number of fractions used, but wrong objectives used for optimisation of remainder of treatment | error made during the scaling of the objectives with the given and to-be-given number of fractions. | sub-optimal treatment plan to catastrophic |  | 1 | 3 | 9 | 3
|  | the accuared dose of tissues can not directly be visualised during the optimasation and compared with known Tolerance doses | sub-optimal treatment plan to catastrophic |  | 1 | 3 | 9 | 3
Plan control
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Position verification
Positioning | intrafraction motion is enhanced with IMRT because patient is longer on the treatment couch. | more MU are given, resulting in longer treatment time | larger margins applied
| using multiple PTV's and cone beam positioning in a single patient, there can be a mix-up of PTV's and correct isocentre | more than one PTV in combination with poor attention | Poor dose coverage of PTV at larger distances from ref point |  | 2 | 3 | 6 | 6 | yes |  |  | yes
disagreement between laser / field indicator and radiationfield
Positioning or fixation aids | different aids used than was intended | protocol violation | treatment delay |  | 4 | 2 | 6 | 8 | yes |  |  | yes
|  |  | sub-optimal treatment |  | 4 | 2 | 6 | 8 | yes |  |  | yes
|  | was not mentioned in note | treatment delay |  | 2 | 2 | 6 | 4
|  |  | sub-optimal treatment |  | 2 | 2 | 6 | 4
ConeBeam od EPID system not propperly calibrated | collision with couch | mis-alignment patient
| not carried out | mis-alignment patient
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Treatment
Treatment | correction for table absorption not valid for all segments | beam apparture at start is not representative for rest of segments | dose delivery variations of few % |  | 3 | 2 | 8 | 6
| intermachines differences limiting the number of fractions to be delivered on an other machine than intended | differences in characteristics regarding dosimerty and mechanics or aids | sub-optimal treatment plan |  | 2 | 2 | 7 | 4
| table position is not restored after table/gantry rotation | protocol violation | higher NTCP or lower TCP |  | 2 | 3 | 3 | 6
| beams transit through metal bars on couch | protocol violation | higher NTCP or lower TCP |  | 2 | 3 | 4 | 6
| difference in planned and delivered MU | beam inhibits may distrub the MU count in the R&V system | Less dose is delivered than planned |  | 1 | 2 | 4 | 2
| too many fractions on different accelerator than planned | not obeying protocols | sub-optimal treatment plan |  | 3 | 2 | 9 | 6
| plan is not accurate anymore | patient geometry has changed sgnificantly | sub-optimal treatment plan |  | 2 | 2 | 8 | 4
|  | plan was accidently changed | sub-optimal treatment plan to catastrophic |  | 2 | 3 | 9 | 6
Workflow
protocols have been altered
software updates
software upgrades
network problems
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Secondary tumours
Treatment | less effective than expected | higher contribution of scattered low energy photons and dose to normal tissue | incidence of RT induced malignancies is doubled | Hall and Wuu
Overtreatment
IMRT becomes standard | IMRT is cost inefficient | application of IMRT for sites where complications are not a limiting factor for the prescription dose | high costs
Dose verification
FILM
Film scanning | analysis is not possible | scanner/computer are missing/broken | treatment delay or sub-optimal treatment |  | 1 | 2 | 1 | 2
|  | software problems | treatment delay or sub-optimal treatment |  | 1 | 2 | 7 | 2
|  | incorrect settings | treatment delay or sub-optimal treatment |  | 1 | 2 | 7 | 2
|  | film is not produced anymore |  |  | 1 | 2 | 1 | 2
| analysis is not accurate | Film contains artifacts, such that quality of treatment cannot be verified accurately |  |  | 4 | 2 | 1 | 8
|  | position of film on scanner is uncertain | treatment delay or sub-optimal treatment |  | 2 | 2 | 1 | 4
|  | film characteristics have changed dramatically |  |  | 2 | 2 | 7 | 4
|  | film calibration is not correct (background, lateral correction) |  |  | 2 | 2 | 7 | 4
Ionisation chamber | analysis is not accurate | non-accurate or expired calibration of IC | treatment delay or sub-optimal treatment
|  | high voltage not switched on | inaccurate or wrong dose
incorrect entry of PT correction
Chamber is too large for fields
for too many beams the IC lies in the penumbra region
2D means of measurement | analysis is not possible | hardware failure |  |  | 3 | 2 | 4 | 6
|  | computer is missing/broken |  |  | 1 | 2 | 1 | 2
|  | faullty software updates/upgrades |  |  | 1 | 2 | 9 | 2
| analysis is not accurate | steep dose gradients compared to detector spacing |  |  | 4 | 2 | 8 | 8
|  | wrong calibration file |  |  | 2 | 3 | 3 | 6
|  | hardware failure |  |  | 2 | 2 | 5 | 4
|  | wrong pressure temperature correction |  |  | 2 | 2 | 9 | 4
|  | corrections for direction sensitivity of equipement fails |  |  | 4 | 3 | 6 | 12 |  |  |  | yes | Eliminate | Verify home-made programmes that interact with clinical plans
|  | faullty software updates/upgrades |  |  | 2 | 2 | 6 | 4
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written
Quality Assuance
Patient specific quality assurance | poor agreement between measurments and calculations | unnotified adjustment in protocol | non optimal treatment or treatment delay |  | 3 | 3 | 2 | 9 |  |  |  | yes | Control | Freeze protocols. Modifications in protocols or procesess may only be carried out after evaluation of it's possible consequences downstream in the process chain.
|  | geometrical changes in patient with in-vivo measurments | non optimal treatment or treatment delay |  | 3 | 3 | 2 | 9 |  |  |  | yes | Control | Regular check of portal images and/or patient weight, or perform in vivo disimetry.
|  | plan was non-typical; assumptions on which dose verification protocols are based were not valid. | non optimal treatment or treatment delay |  | 2 | 2 | 2 | 4
|  | Structures with density overrides were not removed during planning |  |  | 3 | 2 | 4 | 6
|  | ISOC was not placed correctly during planning |  |  | 4 | 2 | 2 | 8
|  | dose grid not placed correctly during planning |  |  | 2 | 2 | 2 | 4
|  | scanningcouch was not removed from images during planning |  |  | 2 | 2 | 2 | 4
|  | incorrect entry of analysis parameters (DTA; Cutoff level, mm, %, ref dose)) |  |  | 3 | 3 | 5 | 9 |  |  |  | yes | Control | Use standard settings and always verifiy.
| good agreement between measurements and calculations is based on a false positive | incorrect entry of analysis parameters (DTA; Cutoff level, mm, %, ref dose))
| plan for QA is not treatment plan | plan was altered after measurements |  |  | 3 | 3 | 5 | 9 |  |  |  | yes | control | Monitor changes in treatment plan.
|  | many plans were available for QA |  |  | 3 | 3 | 6 | 9 |  |  |  | yes | eliminate | Adjust workinstrunctions: remove obsolete plans during plannning after selection of clinical treatment plan.
| measurements have not been performed | not filled out in to-do list | treatment delay |  | 3 | 2 | 6 | 6
|  | to-do list not reviewed |  |  | 3 | 2 | 4 | 6
QA on accelerator | quality is degraded | adjustment of one machine parameters affects other parameter as well (bending-fine vs symmetry) | increased toxicity and decreased tumour control | [6 empty] | yes
|  | no QA has been performed | increased toxicity and decreased tumour control |  | 2 | 3 | 5 | 6 |  | yes
|  | measurements not perforemed according to protocol (not correct equipment, level not checked, high voltage not on, ….) | increased toxicity and decreased tumour control |  | 2 | 2 | 7 | 4 |  | yes
|  | adjusted settings not safed, not verification measuremtns performed. | increased toxicity and decreased tumour control |  | 2 | 3 | 4 | 6 |  | yes
MLC position verification | beyond tolerance | failure in MLC tracking system (laser/lightsource, camera, frame-grabber) | non optimal treatment or treatment delay | Nelms, Sastre-padro | 2 | 2 | 8 | 4 |  | yes
|  | error or inaccuracy in calibration | non optimal treatment or treatment delay | Nelms, Sastre-padro | 2 | 2 | 4 | 4 |  | yes
due to IMRT geen aansluitingen meer die zeer kritiek zijn, Sequencer moet wel iets leveren dat lijkt op Field in Field. Op die manier, in combinatie met meerdere bundelrichtingen en coorilatorhoeken, smeren dosis effecten ten gevolge van onnauwkeurigheid in MLC positionering uit over een groter volume.
| analysis is not possible | wrong software version |  |  | 1 | 3 | 4 | 3
|  | wrong software settings |  |  | 1 | 3 | 4 | 3
|  | film characteristics have changed |  |  | 2 | 2 | 3 | 4
| analysis is not accurate | wrong software version |  |  | 3 | 2 | 6 | 6
|  | wrong software settings |  |  | 3 | 2 | 6 | 6
|  | film characteristics have changed |  |  | 2 | 2 | 3 | 4
MLC speed | leaf lags | wearing of MLC motors | dose distribution deviated from intention |  | 2 | 3 | 2 | 6
End to end test
Output variation | output deviation | ionisation chamber is damaged | large number of patients is irradiated to different dose |  | 1 | 2 | 6 | 2
|  | wrong settings used, not verified |  |  | 1 | 2 | 8 | 2
Fieldsize | asymmetric fields | emphasis on fieldsize (X1+X2) measurements rather than symmetry X1 and X2 | shift in field position, but accurate size
Beamflatness
Info transition | delivered treatment deviates from intended treatment | poor informationtransfer to next planner Orally
poor informationtransfer to next planner Digitally
poor informationtransfer to next planner Written

## Rating Scales RPN
| Continue at: | 9
| Tresholds: | 4 | 8 | 12
| 4 | 3 | 2 | 1
4 | 16 | 12 | 8 | 4
3 | 12 | 9 | 6 | 3
2 | 8 | 6 | 4 | 2
1 | 4 | 3 | 2 | 1